## Slide 1
VOCATIONAL TRAINING REPORT PRESENTATION | B.TECH CSE (AI)
HemoVision
Contactless Vital-Sign Estimation & Biomarker Analysis via Remote Photoplethysmography (rPPG)
SUBMITTED BY
Harsh Patel
Roll No: 300111323021 Enrolment: CD8812 B.Tech 6th Sem, CSE (AI)
SUBMITTED TO
Bhilai Institute of Technology, Durg
Affiliated to CSVTU Bhilai Under Guidance of: Dr. Sumit Sar & Dr. Sunita Soni
TRAINING ORGANIZATION
BS Digital Technology
Vocational Training Period: 08 June 2026 – 08 July 2026 Supervisor: Mr. Kailash Sinha

## Slide 2
01. ABSTRACT & OVERVIEW
Executive Summary & Project Highlights
9.19 BPM
Held-Out Test MAE
523 clips / 91 held-out subjects
598 Subjects
MCD-rPPG Clinical Dataset
3,427 clips across 3 camera types
10 Biomarkers
Feasibility & Confound Audit
Exposed identity shortcuts in RGB AI
• Contactless Physiological Sensing: Measures Blood Volume Pulse (BVP) from ordinary RGB webcam facial video.
• Multi-ROI Geometric Pooling: Tracks 468 3D landmarks via MediaPipe Face Mesh to isolate 8 anatomically stable facial regions.
• AC/DC Signal Disentanglement: Separates high-frequency pulsatile AC signals from static DC melanin/lighting.
• Bounded Dilated TCN: Lightweight 148k-parameter temporal network with bounded tanh waveform supervision.
• 52% Error Reduction: Deep TCN (9.19 BPM MAE) dramatically outperforms classical baselines (POS 19.17 BPM, CHROM 20.28 BPM).
• Ground-Truth Rectification: Discovered and fixed a critical ECG T-wave double-counting bug in post-exercise annotations.
• Rigorous Negative Result: Proved that RGB video cannot measure hemoglobin (r=0.44 was face identity memorization, beaten by demographic OLS r=0.70).
• Live Telemetry UI: Deployed an interactive Streamlit & OpenCV clinical dashboard with real-time waveform playback.

## Slide 3
02. PROBLEM STATEMENT
Problem Identification: Contact Vitals vs. rPPG
• Intrusive Hardware: Requires finger pulse oximeter clips, adhesive ECG chest leads, and inflatable pressure cuffs.
• Patient Discomfort & Tissue Trauma: High risk of skin breakdown in neonatal ICU and elderly/burn patients.
• Infection & Contamination Risk: Physical contact increases cross-contamination in epidemic hospital triage.
• Inconvenient Continuous Telemetry: Cumbersome wires restrict patient mobility during sleep and recovery.
• Invasive Lab Testing: Blood chemistry (Hb, HbA1c, lipids) requires frequent venous blood draws and lab turnaround.
• Optical Contactless Sensing: Ambient video captures sub-visual skin micro-color changes caused by heart pulsations.
• Universal Hardware: Operates with standard consumer webcams, smartphones, and telemedicine laptops.
• ⚠️ Literature Flaw 1 (Data Leakage): Many published works mix subject video clips across train/test splits, memorizing face identities.
• ⚠️ Literature Flaw 2 (Unsubstantiated Biomarker Claims): Papers claim webcam blood chemistry without auditing demographic shortcuts.
• HemoVision Goal: Build a leak-free, mathematically verified rPPG pipeline with falsifiable clinical benchmarking.

## Slide 4
03. RPPG PRINCIPLES
Biological Foundations: Blood Volume Pulse Optics
• Cardiac Systole & Diastole: Heart contraction pumps oxygenated blood into facial capillary networks.
• Capillary Expansion: Subcutaneous micro-vessels expand rhythmically with every pulse wave.
• Blood Volume Pulse (BVP): Modulates the optical density and path length of light penetrating skin tissue.
• Green Channel Peak (520–580 nm): Hemoglobin exhibits peak optical absorption in green wavelengths.
• Micro-Color Variations: Produces a tiny intensity fluctuation of ~0.1% to 1.0% of mean skin pixel value.
• Red & Blue Channels: Act as motion and specular reflection references to isolate arterial blood changes.
• DC Component (Static ~99%): Skin melanin, baseline tissue absorption, and constant ambient illumination.
• AC Component (Pulsatile ~1%): Dynamic cardiovascular pulse wave signal.
• Z-Score Filtering: Eliminates DC baseline offsets while keeping pulse shape for neural TCN regression.

## Slide 5
04. CLINICAL DATASET
MCD-rPPG Clinical Dataset Architecture
• 3,427 Total Video Clips across 598 unique subjects.
• Leak-Proof Subject Splits: Train (418 sub, 2,387 clips), Val (89 sub, 517 clips), Test (91 sub, 523 clips). Zero subject overlap.
• 3 Consumer Camera Sensors: FullHDwebcam, USBVideo, IriunWebcam.
• 2 Physiological States: Resting baseline ('before') and Post-exercise stress ('after').
• 11 Ground-Truth Clinical Labels: ECG Heart Rate, Respiratory Rate, SpO2, Systolic/Diastolic BP, Hemoglobin, HbA1c, Cholesterol, Rigidity, Stress, and BMI.

## Slide 6
05. PREPROCESSING
Facial Mesh Tracking & Multi-ROI Preprocessing
• MediaPipe 468 3D Landmarks: Dynamically aligns facial skin masks invariant to head tilt and position.
• 8 Anatomical ROIs: Left/Right Forehead, Left/Right Upper Cheek, Left/Right Lower Cheek, Nose Bridge, and Chin.
• 24-Channel Tensor: Spatial RGB averaging per ROI yields a [8 ROIs × 3 Channels × T Frames] matrix.
• Color Channel Auto-Fix: Validates RGB order via skin reflectance heuristic (R_mean > B_mean).
• Butterworth Bandpass: 2nd-order zero-phase filter (0.65–3.25 Hz = 39–195 BPM) removes slow drifts and high-frequency noise.

## Slide 7
06. CLASSICAL BASELINES
Classical Benchmarks & FFT Quantization Fix
• POS (Plane-Orthogonal-to-Skin): Projects RGB onto skin-tone invariant subspace (MAE = 19.17 BPM on test set).
• CHROM (Chrominance-based): Color-difference projection with bandpass subtraction (MAE = 20.28 BPM).
• 🔍 FFT Quantization Bug Fixed: Standard Welch PSD was locking HR estimates to rigid 7.5 BPM bins (nperseg=240 at 30 FPS -> 0.125 Hz).
• Zero-Padding Solution: Applying N_FFT = 4096 unlocked continuous sub-BPM frequency resolution.
• Honest Finding: Even after zero-padding, classical MAE remained ~18.9 BPM—proving classical projection is fundamentally noise-limited.

## Slide 8
07. NEURAL ARCHITECTURE
HemoVisionBoundedTCN Architecture
• Input Tensor: [B, 24, T] containing 8 ROI × 3 RGB channels over 600 frames (~20s).
• Pointwise 1D Conv: Projects 24 raw channels to 48/64 hidden dimensions.
• Normalization & Activation: BatchNorm1d + GELU activation for smooth non-linear representation.
• Depthwise Separable Convs: Decouples spatial channel mixing from temporal sequencing (78% parameter reduction).
• Exponential Dilation: Rates d = (1, 2, 4, 8) expand receptive field to >20 seconds without pooling.
• Residual Connections: x_out = GELU(Conv(x) + x_res) ensures robust gradient flow.
• Bounded Waveform Head: 1D Conv + tanh bounds BVP output to [-1.0, +1.0], preventing gradient explosions.
• Multi-Head Attention: 4-head self-attention on AC trunk features for global temporal pooling.
• Multi-Task Heads: Regresses pulse waveform alongside 10 tiered clinical biomarker estimates.

## Slide 9
08. MODEL TRAINING
Loss Formulation & Model Progression
• Waveform Loss (Negative Pearson r): Directly maximizes correlation with ground-truth BVP wave: Loss = 1.0 - Pearson_r(y_pred, y_true)
• Scale & Offset Invariant: Avoids penalizing arbitrary amplitude/baseline shifts in optical signals.
• Tier-Weighted Biomarker Loss: Balances established vitals (weight 2.0) vs. exploratory biometrics (weight 0.3).
• Optimization: AdamW (LR=3e-4, weight_decay=1e-4), CosineAnnealingWarmRestarts, and AMP Mixed Precision.
• Gradient Accumulation: Effective batch size of 16 enables stable convergence on a 4GB GPU.

## Slide 10
09. DATA INTEGRITY
Data Integrity: The ECG T-Wave Double Counting Fix
• Anomaly in Post-Exercise Clips: In 'after-exercise' recordings, elevated pulse rates caused elevated T-wave amplitudes in ECG leads.
• Automated Peak Detector Flaw: The clinical ground-truth algorithm erroneously counted large T-waves as secondary R-peaks.
• Doubled Heart Rate: Ground truth HR was inflated by up to 2× (reporting 180+ BPM for true 90 BPM pulse).
• Initial Confusion: All rPPG models appeared to fail catastrophically (MAE ~48 BPM) due to corrupt ground truth.
• Adaptive Threshold Re-calibration: Enforced minimum refractory period (250 ms) and dynamic amplitude thresholds.
• Ground-Truth Rectification: Restored true biological heart rates across all affected post-exercise recordings.
• Immediate Error Drop: Classical baseline error dropped immediately from ~48 BPM to ~19 BPM.
• Unlocked Deep Learning: Enabled clean gradient signals for `HemoVisionBoundedTCN`, reaching 9.19 BPM test MAE.
• Key Takeaway: Real-world AI debugging requires verifying annotation sensors, not just tuning neural weights.

## Slide 11
10. EXPERIMENTAL RESULTS
Experimental Results: Heart Rate Accuracy
• 9.19 BPM Overall MAE across all 523 held-out test clips (91 subjects never seen during training).
• 52% Error Reduction: Slashes error compared to classical POS (19.17 BPM) and CHROM (20.28 BPM).
• High-Quality Demo Clips: Achieves 1.33 BPM MAE on stable lighting recordings (e.g., 101.7 vs. 102.1 BPM).
• 56% of Clips Under 5 BPM Error in uncurated 30-clip random evaluations.
• Robust Across Camera Sensors: Generalizes well across FullHDwebcam, USBVideo, and IriunWebcam inputs.

## Slide 12
11. BIOMARKER AUDIT
Multi-Biomarker Feasibility: Rigorous Negative Result
1. Target Standardization: z-score per split -> No improvement in blood chemistry.
2. BatchNorm Freezing: Freeze BN after 5 epochs -> HR stable, biomarkers unchanged.
3. Rich Multi-Head Pooling: Mean+std+attention -> More confident, not more accurate.
4. Raw DC Bypass: Separate x_raw skin tone path -> No biomarker improvement.
5. Lin's CCC Loss: Concordance loss -> HR worse, biomarkers unchanged.
• Definitive Scientific Conclusion: The bottleneck is physical information content in RGB video, NOT model capacity.

## Slide 13
12. CONFOUND ANALYSIS
Hemoglobin Deep Dive: Exposing Identity Shortcuts
• Apparent Video Correlation: The video model achieved apparent r = 0.442 (MAE 1.24 g/dL).
• Demographic OLS Baseline: A simple linear model using only Age, Sex, & BMI achieved r = 0.700 (MAE 0.95 g/dL).
• The Flaw (exact_repeat_fraction = 1.0): Venous blood was drawn once per subject visit. All clips for a subject share the identical Hb label.
• Face Recognition Shortcut: The neural net learned facial morphology/identity cues rather than optical blood absorption.
• Physics Reality: RGB cameras lack dual-wavelength (660/940 nm) NIR optics required to measure hemoglobin.

## Slide 14
13. SOFTWARE DEPLOYMENT
Interactive Clinical Telemetry Dashboard
• Real-Time Video Ingestion: Supports both local webcam feeds and recorded benchmark video files.
• Dynamic 8-ROI Overlay: Renders MediaPipe facial mesh polygons with live color tracking in OpenCV.
• Motion Stability Monitor: Tracks head displacement to flag motion artifacts in real time.
• Reconstructed BVP Wave: Real-time scrolling plot of neural-predicted pulse waveform.
• Synchronized Ground Truth: Side-by-side comparison against contact ECG/PPG sensor data.
• Signal Quality Index (SQI): Real-time signal-to-noise ratio confidence gauge.
• Welch PSD Frequency Spectrum: Zero-padded FFT peak-picking inside 39–195 BPM band.
• Instantaneous BPM Counter: Large digital gauge with instantaneous and sliding window HR.
• Zero Server Latency: Fully self-contained local inference with cached PyTorch weights.

## Slide 15
14. TRAINING OUTCOMES
Industrial Training Outcomes & Technical Skills
• 468-point 3D landmark localization via MediaPipe.
• Multi-ROI polygon mask rasterization and spatial color pooling.
• Real-time video buffering and frame rate synchronization.
• 2nd-order zero-phase Butterworth bandpass filtering (0.65–3.25 Hz).
• AC/DC signal disentanglement and z-score normalization.
• Classical POS / CHROM projection implementations and Welch PSD.
• Depthwise-separable 1D temporal convolutions with dilation schedules.
• Negative Pearson correlation loss & multi-head attention pooling.
• Mixed precision (AMP) and gradient accumulation on GPU.
• Discovered ECG T-wave doubling & FFT quantization bugs.
• Leak-proof subject-disjoint evaluation and demographic confound audit.
• Real-time Streamlit and OpenCV clinical dashboard deployment.

## Slide 16
15. FUTURE WORK
Recommendations & Future Research Roadmap
• Dual-wavelength active LED illumination (660 nm visible red + 940 nm NIR).
• Provides physical optical absorption spectra required for true non-contact SpO2 and hemoglobin.
• Global-shutter 60–120 FPS sensors to eliminate rolling-shutter distortion.
• Inter-Beat Intervals (IBI) for autonomic stress markers (RMSSD, SDNN, LF/HF).
• Pulse Wave Transit Time (PWTT) via dual-site ROIs (forehead vs. chin delay) for continuous cuffless blood pressure.
• Morphological dicrotic notch analysis.
• Model quantization to ONNX / TensorRT INT8.
• Low-power deployment on edge embedded hardware (Raspberry Pi 5, NVIDIA Jetson).
• Lightweight web-assembly (Wasm) in-browser execution.
• FHIR / HL7 standard REST API integration for electronic health records (EHR).
• Real-time automated patient triage alerts (tachycardia, bradycardia, arrhythmia).
• Multi-patient simultaneous tracking in hospital waiting rooms.

## Slide 17
16. CONCLUSION
Conclusion & Key Takeaways
1. Heart Rate is Fully Measurable via Webcams
Achieved 9.19 BPM MAE on 91 held-out subjects (523 clips) and 1.33 BPM MAE on demo clips, cutting classical errors in half.
2. Rigorous Negative Result for 10 Biomarkers
Proved that ambient RGB video cannot measure blood chemistry (Hb, HbA1c, lipids); exposed apparent correlation as face identity memorization.
3. Engineering Integrity & Hidden Bug Rectification
Discovered and fixed the ECG T-wave doubling anomaly and FFT quantization locking, establishing leak-proof evaluation protocols.
4. Full-Stack Clinical Telemetry Software
Constructed an interactive real-time Streamlit and OpenCV dashboard displaying synchronized facial ROIs, live BVP waveforms, and Welch PSD.

## Slide 18
VOCATIONAL TRAINING REPORT PRESENTATION
Thank You! Questions & Answers
Key Academic References:
1. Wang, W., et al. (2017). Algorithmic principles of remote photoplethysmography. IEEE TBME.
2. de Haan, G., & Jeanne, V. (2013). Robust pulse rate from chrominance-based rPPG. IEEE TBME.
3. Lugaresi, C., et al. (2019). MediaPipe: A Framework for Building Perception Pipelines. arXiv.
4. Paszke, A., et al. (2019). PyTorch: An Imperative Style Deep Learning Library. NeurIPS.
Harsh Patel | Roll No: 300111323021 | B.Tech CSE (AI), Bhilai Institute of Technology, Durg Vocational Training at BS Digital Technology (June – July 2026)